## SW Info
*SW Info Tab may be filled in on a separate document, if desired, as long as content is maintained
Plan/Report Identification
| State | MWS DOC# | Rev
| Plan | DOC2053697 | 1
| Report | DOC2053697 | 2
| Plan | DOC2053697 | 3
| Report | DOC2053697 | 4
| Plan | DOC2053697 | 5
Document Revision History
Complete the table below for each revision. Add lines as necessary. Revision number should correspond to electronic revision number.
| Rev # | Author | Summary of Change
| 1 | Ralph Williams | Software Validation Plan
| 2 | Ralph Williams | Software Validation Report
| 3 | Sirisha Kota | Software Validation Plan
| 4 | Brad Peterson | Software Validation Report
| 5 | Sirisha Kota | Software Validation Plan
Software Info
| Equipment ID |  | Windows 7 or equivalent
| Software File Name & Version: |  | ZeissPDI.jar
| Software Archive Location |  | DOC2054880
| Configuration (PDL/Recipe) Documents |  | 5781640PDL
Operating Environment including any required supporting or interfacing hardware and software (including configurations, versions, utilities, fixtures, etc)
Hardware: PC Software: Microsoft Windows 7 or equivalent Java8 JRE (1.8.XXX)
Software Purpose & Intent
Description of what software does, software architecture, flow chart, etc..
This software validation is for the MercuryG "ZeissPDI" (Pack2.0, Single Pack) user interface. This software will copy the raw data output from the Zeiss O-Select camera into virtual memory, compile the raw data, perform statistical analysis to determine Pass/Fail specs on the data, and finally display the Pass/Fail grade to the User. When all relevant data is collected, the software will store the complete records into the database. A complete record consists of the associated compiled data records, a Pass/Fail grade for each spec, as well as an overall Pass/Fail from grade each of the four measured MercuryG Packs contained on the Frame that was used during the dicing process. Revision 2: This revision of the software is to accommodate the following requirements: 1. Minimize the lag to SQL server 2. Remove the restriction to fully populate the frame 3. Write one pack at a time to SQL (instead of 4) 4. Calculate delta from nominal value instead of the actual measurement value Revision 3: This revision of the software is to fix the following bug: The UI was not writing data to SQL for packs which were identified as Visual Defects.
Software Type
Commercial Off the Shelf (COTS)
Configurable Software
Custom Built Software
Software Impact
| Major | Failure of the automated equipment impacts decisions that affect product quality. (e.g. product acceptance criteria, test data analysis).
| Minor | Failure of the automated equipment impacts quality system record integrity (e.g., record loss, record corruption).
Severity Justification: add text here
This software will be used for acceptance testing of the MercuryG Pack2.0/Single Pack product at the Post-Dice Inspection process.
| High | A failure of the system will likely be detected through downstream controls (e.g. process or product controls, manual inspection, etc)
| Low | A failure of the system is unlikely to be detected through downstream controls
Detectability Justification: add text here
100% downstream verification
Detectability
|  |  | High | Low
| Severity | Minor | Low | Medium
|  | Major | Medium | High
Software Impact (High, Medium, or Low):
Medium See 5111000QPA
Software Installation
Define how and where the software is loaded onto the equipment.
i.e. folder file path or special software needed (PLC Proficy Machine Edition)
Place correct version of the ZeissPDI.jar file on the PC.
Software Validation Conclusions (REPORT STATE)
Yes
No
Yes
No
Yes
No
Validation Summary
Software Validation Statement
Validation Result:
Pass
Fail
Name:
Date:
Required Approvers
Approval for the Plan and Report State shall be in MyWorkshop Routes. If approval route is independent of equipment qualification route, then route to the roles below. Additional approvers and reviewers are acceptable.
| Role |  | Name
| Engineering Rep/Product Champion |  | Mike Prescott
| Manufacturing Project Leader or Senior Manufacturing Engineer or Equipment Owner (not the author or tester) |  | Nick Lodl
| Manufacturing/Test Equipment Engineer or designee (Author) |  | Brad Peterson
| Quality Assurance |  | Tony Gillon
| Tester (may be the author) |  | Brad Peterson
| Other (if necessary) |  | NA
Identify the SEVERITY of software based on the following criteria:
Identify the DETECTABILITY of software failure based on the following criteria:
If Software Impact is Medium or High, document the mitigations in your process or equipment that reduces the risk:
Select the Software Type
Is the most recent version of the software archived in the location listed above?
Is the most recent version of the recipe/program/PDL archived in the location listed above?
Deviations (Fill out deviation attachments 2 & 3 and include with Report). Were there deviations?

## Requirements
Must be filled in at the Plan State
REQ # | Requirement Add or Modify Requirements specific to automated equipment under test
1 | Operating Environment
1.1 | All related hardware and software can be installed and functionally interfaces.
2 | Software Name, Version & Start Up
2.1 | Proper version of ZeissMeasure UI is installed.
3 | Security and User Permissions
3.1 | Settings window password protected
4 | Recipe\Routine
4.1 | NA - There are no recipes or routines.
5 | Alarms, Sensors, Error Codes & Interlocks
5.1 | Zeiss connectivity check
5.2 | Database connectivity
6 | Software Functional Tests
6.1 | ZeissMeasure software will start the Ziess O-SELECT.
6.2 | UI Settings
6.3 | Queue/Frame population
6.4 | Data compilation
6.5 | Display results
6.6 | MaxSize Dialog
6.7 | Defect Manager
6.8 | Remove the restriction to fully populate the frame
6.9 | The UI should write data to SQL for packs which are identified as Visual Defects. | Bug
7 | Overall Software Functionality
7.1 | Overall software functionality for Merc-G Pack2.0 Post-Dice
7.2 | Overall software functionality for Merc-G Pack2.0 Pixel-Measure
7.3 | Overall software functionality for Merc-G Single Pack Post-Dice
7.4 | Overall software functionality for Merc-G Single Pack Pixel-Measure

## Test Scripts
|  | REQ # | User Requirement | Test Procedure / Method | Expected Results | Actual Results | Pass / Fail | Verifier (Name / Date)
Requirement 1: Operating Environment
1 | 1 | 1.1 | Verify all related hardware and software can be installed and functionally interfaces. | See test cases for section 2.1. | See test cases for section 2.1. | NA | NA | NA
Requirement 2: Software Name, Version & Start Up
2 | 1 | 2.1 | Proper version of ZeissMeasure UI is installed. | Under File-About, the proper Version is listed. | The GUI shows: Filename:Zeiss Post-Dice User Interface Version 1.0.1
Requirement 3: Security and User Permissions
3 | 1 | 3.1 | Settings window properly password protected. | Under File-Settings Enter the password engjjs750 into the password text box and push enter. | Settings window opens properly | Settings window opened properly when password was entered correctly. | PASS | Ralph Williams 2017-Oct-23
Requirement 4: Recipe\Routine
4 | 1 | 4.1 | N/A | N/A | N/A
Requirement 5: Alarms, Sensors, Error Codes & Interlocks
5 | 1 | 5.1 | Zeiss connectivity check | Change the setting Ziess Window Name to something ofther than the actual window name. User attemps to measure part. | The UI will display message that the Zeiss O-SELECT - '{BAD WINDOW NAME}-'is not running | UI displayed message "Zeiss O-SELECT - MercG_Pack2_WSBlank_Rev3' is not running" | PASS | Ralph Williams 2017-Oct-23
5 | 2 | 5.2 | Database connectivity | Change the server IP address to 0.0.0.0 in settings tab, and attempt to run the software. | UI indicates that a connection to the database could not be made | UI displayed "There was a problem writing csv. A connection to the database could not be made!" | PASS | Ralph Williams 2017-Oct-23
Requirement 6: Software Functional Tests
6 | 1 | 6.1 | ZeissMeasure software will start the Ziess O-SELECT. | Click measure with Zeiss O-SELECT window open. | The Zeiss O-SELECT will take a measurement | The Zeiss O-SELECT takes a measurement. | PASS | Ralph Williams 2017-Oct-23
6 | 2 | 6.2 | UI Settings | In settings window, there is a modifiable field to change the Outlier Multiplier. Open the Settings page (File - Settings), enter password, change the Outlier Multiplier to desired value, save and close the Settings page. | The Outlier Multiplier will be updated and is applied when beginning the next full measurement. | The Outlier Multiplier is applied as expected | PASS | Ralph Williams 2017-Oct-23
6 | 3 | 6.3 | UI Settings | In settings window, there are modifiable fields to change individual spec limits, as well as the order of priority to which the specs are applied to the Pack. Hold a tech review with process owner to ensure modified spec limits and the order of priority are applied correctly to the processed data. | Modified spec limits and the order of priority are applied correctly to the processed data. | Modified spec limits and the order of priority was applied correctly to the processed data. | PASS | Ralph Williams 2017-Oct-23
6 | 4 | 6.4 | Queue/Frame Population | Click Populate Frame button, enter good data into required fields. User selects a Dicer and populates the frame with the individual pack ids that were Diced on the Frame. | A queue of the PackID's entered are displayed in the Current Frame list. | A queue of the PackID's entered are displayed in the Current Frame list. | PASS | Ralph Williams 2017-Oct-23
6 | 5 | 6.5 | Queue/Frame Population | Click Populate Frame button, enter Bad/Missing data into required fields | User presented with error dialog indicating there was a "Problem creating the Frame" | User presented with error dialog indicating there was a "Problem creating the Frame" | PASS | Ralph Williams 2017-Oct-23
6 | 6 | 6.6 | Data compilation | Hold a tech review with process owner to ensure math is applied correctly to the Zeiss O-SELECT raw data output for Pixel-Measure, before analysis of acceptance criteria is performed. | The math is applied correctly to the Pixel-Measure raw data. | Math applied correctly to the Pixel-Measure raw data. | PASS | Ralph Williams 2017-Oct-23
6 | 7 | 6.7 | Data compilation | Hold a tech review with process owner to ensure math is applied correctly to the Zeiss O-SELECT raw data output for Post-Dice, before analysis of acceptance criteria is performed. | The math is applied correctly to the Post-Dice raw data. | Math applied correctly to the Post-Dice raw data. | PASS | Ralph Williams 2017-Oct-23
6 | 8 | 6.8 | Display results | Process the file Merc_Combined_PDI__PM_Rev3_chr.txt that is known to PASS all DOC specs | The software will display the results with a 'PASS' indicator | Software displayed 'PASS' indicator when expected | PASS | Ralph Williams 2017-Oct-23
6 | 9 | 6.9 | Display results | Process the file Merc_Combined_PDI__PM_Rev3_chr.txt that is known to FAIL one or more DOC specs | The software will display the results with a 'FAIL' indicator | Software displayed 'FAIL' indicator when expected | PASS | Ralph Williams 2017-Oct-23
6 | 10 | 6.1 | Display results | Process the file Merc_Combined_PDI__PM_Rev3_chr.txt that is known to FAIL the DOC spec as an OUTLIER | The software will display the results with a 'REMEASURE' indicator | Software displayed 'REMEASURE' indicator when measurement exceeded OUTLIER | PASS | Ralph Williams 2017-Oct-23
6 | 11 | 6.11 | MaxSize dialog | Process the file Merc_Combined_PDI__PM_Rev3_chr.txt, once the raw data is compiled, the DEFAULT Pass/Fail MaxSize dialog is presented to the user. | The user selects Pass/Fail as determined when inspecting the Pack with MaxSize Go/NoGo gauge. The MaxSize spec will be Pass/Fail as selected. | Selected 'Pass', MaxSize spec Passed. Selected Fail', MaxSize specFailed. | PASS | Ralph Williams 2017-Oct-23
6 | 12 | 6.12 | Defect Manager | Press the 'Defect Manager' button, select the appropriate defect from the list, click the 'Add Defect" button. | The Defect Manager dialog is presented to the user, the appropriate defect is selected from the list. The current Pack FAILS and displays 'FAIL' indicator. The FAIL_CODE field is populated with the defect. | Defect Manager dialog was displayed, defect was selected from the list, the current Pack FAILED and displayed the 'FAIL' indicator. The FAIL_CODE field was populated with the defect. | PASS | Ralph Williams 2017-Oct-23
6 | 13 | 6.13 | Remove the restriction to fully populate the frame | Click Populate Frame button, enter Bad/Missing data into required fields. | A list of the PackID's entered are displayed in the Current Frame list. | Limit removed. Any number of packs can be entered into frame. | PASS | Brad Peterson 2017-Dec-04
6 | 14 | 6.14 | The UI should write data to SQL for packs which are identified as Visual Defects. | Click the "Defect Manager" button Select the type as "Visual" Click the "Add Defect" button | The data should be written to SQL for packs which are identified as Visual Defects. |  |  |  | Bug
Requirement 7: Overall Software Functionality
7 | 1 | 7.1 | Overall software functionality for Merc-G Pack2.0 Post-Dice | Perform Post-Dice measurements on all Packs in the "Frame Queue" for MercG Pack 2.0 | The UI will process the Post-Dice data of each individual MercG Pack2.0 Pack, giving a PASS/FAIL result and overall FAIL_CODE for each Pack in the "Frame Queue", as determined by the spec priority. Once each Pack in the "Frame Queue" has a Pass/FAIL result, the completed frame data for each Pack is written to the respective MERCURYG.PDI database
7 | 2 | 7.2 | Overall software functionality for Merc-G Pack2.0 Pixel-Measure | Perform Pixel-Measure measurements on all Packs in the "Frame Queue" for MercG Pack 2.0 | The UI will process the Pixel-Measure data of each individual MercG Pack2.0 Pack giving a PASS/FAIL result and overall FAIL_CODE for each MercG Pack2.0 Pack Pack in the "Frame Queue", as determined by the spec priority. Once each Pack in the "Frame Queue" has a Pass/FAIL result, the completed frame data for each Pack is written to the respective MERCURYG.PIXELMEASURE database.
7 | 3 | 7.3 | Overall software functionality for Merc-G Single Pack Post-Dice | Perform Post-Dice measurements on all Packs in the "Frame Queue" for MercG Single Pack | The UI will process the Post-Dice data of each individual MercG Single Pack, giving a PASS/FAIL result and overall FAIL_CODE for each Pack in the "Frame Queue", as determined by the spec priority. Once each Pack in the "Frame Queue" has a Pass/FAIL result, the completed frame data for each Pack is written to the respective MERCURYG.PDI database
7 | 4 | 7.4 | Overall software functionality for Merc-G Single Pack Pixel-Measure | Perform Pixel-Measure measurements on all Packs in the "Frame Queue" for MercG Single Pack | The UI will process the Pixel-Measure data of each individual MercG Single Pack giving a PASS/FAIL result and overall FAIL_CODE for each MercG Pack2.0 Pack Pack in the "Frame Queue", as determined by the spec priority. Once each Pack in the "Frame Queue" has a Pass/FAIL result, the completed frame data for each Pack is written to the respective MERCURYG.PIXELMEASURE database.

## Objective Evidence
2.1
3.1
3.1
5.1
5.2
6.3.1
6.3.1
6.3.2
6.1
6.4.3
6.4.2
6.4.4
6.5
6.6